# Improvement of Spatial Chromaticity Measurement of white LED luminaires

Jianping Wang1, Yiqing Liu2, Wenpeng Wang1

Zhejiang University Sensing Instruments Co., Ltd, Hangzhou, China

Xiamen Products Quality Inspection Institute, China

The application of white LEDs is wider and wider in lighting, with the improvement of their luminous efficiency. However, the non-uniformity of spatial chromaticity is one of the important issues to white LED luminaires. The chromaticity non-uniformity might be caused by the non-uniformity of phosphor layer, effective optical depth inconsistence of blue beam excitation from chip in phosphor layer, dispersion of attached secondary optics, and interference of the coating which is used to improve the reflecting and transparent performance of optical surface.

There are two methods for spatial chromaticity measurement of LED luminaires: a colorimeter and spectroradiometer to be combined with the goniophotometer. A calibrated CCD spectroradiometer is generally used in the goniophotometer to get absolute spatial chromaticity. However, for low intensities, the exposure time of CCD and data transport limits the measuring speed. It will cost lots of time to measure the non-uniformity of the spatial chromaticity of LED luminaires. The advantages of broadband colorimetric measurement are quick response, good repeatability and accurate measurements of chromaticity coordinates for thermal sources. However, the main disbenefit of those measurement which has to be concerned is colour mismatch of the detectors.

For white LED luminaires used in lighting, although chromaticity might be some variation in space, the spectral distribution is generally very similar. Low cost CCD spectroradiometers could be used to measure the spectral distribution on the optical axis of the LED luminaire of interest (generally in normal). The colour mismatch of the detectors with the spectral response sxr(λ), sxb(λ), sY(λ), sZ(λ) can be corrected based on the above spectral distribution, and the  accuracy for the measurement of spatial chromaticity of LED luminaires could be improved.

In this way, the spatial chromaticity of both LED PAR lamp and LED luminaire is investigated. The chromaticity values in several selected directions were measured by a high accuracy scanning spectroradiometer, and the colorimeter in the goniophotometer at this work is used to determine non-uniformity of spatial chromaticity. And then the measuring results of colorimeter could be corrected with the spectral distribution, which is measured by a portable CCD spectroradiometer with 50 mm focus length. The results show that the uncertainty of absolute chromaticity coordinates could be reduced from 0.02 for x, y to 0.0015 for x, y.

In conclusion, it’s possible that the measurement of chromaticity of LED luminaires in full space could be implemented quickly in the colorimetric goniophotometer by using low cost spectroradiometers.

Keywords: Chrometricity measurement, LED, goniophotometer.